**D2 Daily X-ray Safety checks**

Assuming instrument and computer are up and running:

1. Check to see if instrument sample compartment is clean
2. Check that all vents free from obstruction
3. Check that the rear coolant level indicator is above minimum level
4. Verify that all three dosimeters are in place
5. Check to see if top LED's come on when X-rays are turned on via the software
6. Check that door does not open when X-rays and top LED bar is on